How do I install a Pre-Calibrated Alcohol Sensor Module?

For the AL6000 please take the following steps to install a new sensor module 

(Click on the images for larger versions):

| Prepare your AlcoMate Prestige. | Use a flat, sharp tool to carefully open the top cover. | Lift the breath tube to access the sensor module. |
| --- | --- | --- |
| Remove the breath tube. It should remove easily. | Carefully raise the sensor module from the body. Do NOT detach the cable from the body. | Remove the cable from the slot underneath the sensor module. |
| Prepare a new pre-calibrated sensor module. | Insert the cable into the new sensor. Make sure the blue strip is toward the sensor, as shown. | Replace the sensor module into its original position. |
| Re-attach the breath tube directly on to the sensor. Make sure the front opening is aligned. | Re-attach the cover on to the top of the body. It should snap in fairly easily. | Ready for Testing! |