Electronic Digital Thermometer Operation Instructions MT-B231

### FEATURES

### 1. Internal function test will be carried out every time when the unit is turned on. It is so-called automatic self-test function. If it displays , it means the thermometer is not working properly. Please contact the supplier.

### 2. Temperatures can be taken from under an arm (external use) or under the tongue (oral use). It is suggested that rectal temperatures are taken on infants only when external use is inconvenient.

3. This is a very sensitive device and results can be obtained within 60 ~ 90 seconds approximately.

4. If the thermometer is inadvertently left on after reading your temperature, it will automatically shut off after 10 minutes.

5. This is a small and light-weight unit (10 grams) and can be used by the whole family.

6. The ABS resin body presents no hazards compared to a broken mercury glass thermometer.

### PARTS DESCRIPTION

### TECHNICAL SPECIFICATIONS

| Measurement range: | 89.60 °F ~ 107.60 °F |
| --- | --- |
| Accuracy: | Fahrenheit: ±0.30 °F Less than 95.90 °F ±0.20 °F 95.90 °F ~ 107.60 °F ±0.30 °F 107.60 °F ~ 109.40 °F |
| Display unit: | 0.10 °F |
| Battery: | One 1.55V button size battery (LR 41) |
| Power consumption: | 0.15 MW in measurement mode |
| Battery life: | More than 200 hours of continuous operation |
| Weight: | Approx. 10 grams including battery |
| Dimensions: | 132.3mm x 20mm x 14mm |
| Storage temperature: | -13 °F ~ 131 °F |
| Operation temperature: | 50 °F ~ 95 °F |
| Max. Storage Humidity: | 80% ~ 90% |
| Type: | Non-predictive |
| The digital thermometer is waterproof. |

### HOW TO REPLACE THE BATTERY

When the battery is almost exhausted, the battery symbol “" will appear on the LCD.

● Slide back the battery cover.

● Use a pointed object to remove the battery from its holder carefully.

● Insert a new battery (LR41) with the “+” symbol at the top.

● when replacing the battery cover, take care that the seal is not damaged.

#### NOTE: Batteries should not be disposed of with household waste. Please follow local battery disposal regulations.

WARNING

1. Do not swallow the battery which may be fatal.

2. Do not recharge the battery.

3. Do not dispose of the battery in fire as it may explode.

4. Do not drop the unit.

5. Do not boil the sensor, instead clean the unit by wiping it with dry cloth and clean the sensor with medical alcohol (95%).

6. Do not expose the unit to direct sunlight or high temperatures.

7. Do not bend the unit out of shape.

8. Do not attempt to disassemble the unit except battery replacement.

9. Do not bite or pull the unit.

10. When inserting the unit (for oral and rectal use), insert it gently. Do not insert it strongly or bend it.

HOW TO OPERATE

- Press button (, and the LCD screen  will display 188.88  for approximately 2 seconds. At the same time, an “LCD test” (to check whether all segments of the display appear) and an internal function test are carried out.

- After switching on the thermometer the display will show the last memorized data for approx. 2 seconds.

- The indicator will then display “Lo” for about 0.6 seconds.

- While the temperature is being measured, the current temperature will be permanently indicated and the “°F” mark will flicker during the measuring time.

- When the fluctuation of the temperature is below 0.04°F for 8 seconds, the “°F” mark will stop flickering and an electronic buzzer will sound. At the same time the final temperature will be shown on the LCD. From now on the indicated temperature will fluctuate minimally.

- After measuring your temperature the temperature will be displayed on the LCD screen until you switch off the thermometer.

- To prolong the life of the battery, please switch off the thermometer after use by pressing button (. If the power is left on, the thermometer will automatically shut down after 10 minutes to save power.

NOTE1: When measuring the temperature under your arm, it is recommended that keeping the thermometer in position for 1 to 2 minutes. The buzzer will sound several times during measuring time.

NOTE2: When the body temperature reaches 100.04°F, the fever alarm will sound.

RELIABLE METHOD TO MEASURE TEMPERATURE

Always measure your body temperature in the same place and at the same time of day.

| Method | Measuring Time | Normal Temperature Range | Remark |
| --- | --- | --- | --- |
| ORAL | 1~2 minutes approx. | 95.90°F ~ 99.50°F | Place the sensor under tongue. Sensor tip should touch the root of tongue, as far to the end as possible. Hold the thermometer in the same spot with your mouth closed. Do not shift the thermometer during measurement. |
| EXTERNAL | 1~2 minutes approx. | 94.46°F ~ 99.14°F | Although simpler, the external method is less accurate and takes longer. Before taking your temperature, make sure the armpit is dry. Place the thermometer into the person's armpit with sensor tip pointing upward. Firmly fold the arm to hold the thermometer in place. |
| RECTAL | 1 minute approx. | 97.88°F ~ 100.40°F | This is the most reliable measurement method. Lubricate the probe with a water soluble jelly for easier insertion. Do not use petroleum jelly. Gently slide the probe about 1 cm (less than 1/2 inch) into the rectum. If you detect resistance of any kind, STOP. |

THERMOMETER CLEANING SOLUTION

When continuous re-measurements are being done clean the unit with medical alcohol (95%), then wipe it using a clean dry cloth or dry medical cotton, before you use the unit.

EXPLANATION OF LOT NO.

The date of manufacture is included in the LOT. The LOT has 5 numbers. The first number means the year (0-9). The second number means the month (Jan. - Sep. : 1-9, Oct.: A, Nov.: B, Dec.: C). The third and fourth numbers mean the date (1-31). And the last number means the batch of the month.

DISPOSAL

- Please dispose of the product and the battery separately.

- The product and the battery should not be disposed of with household waste.

- Please follow local waste / battery disposal regulations.

LIMITED WARRANTY

This thermometer is guaranteed for 2 years from the date of purchase against Manufacturer's defect under normal use. If your unit does not function properly due to defective parts or assembly, we will repair it free of charge. All parts are covered by this warranty excluding the battery. The warranty does not cover damage to your unit due to improper handling.

Applying standards: 93/42/EEC MDD, IEC60601-1:1990, IEC60601-1-2:2001, EN12470-3:2000, ASTM E1112:98